## Datasheet
JM Test Systems, Inc. |  |  | Calibration Data Sheet
Function / Range |  |  | Out | Measured Value |  | Acceptance Limits
Test Description | Nominal Value |  | * | AS FOUND | AS LEFT | LOW |  | HIGH
TEST RESULT |  |  |  | Pass | Pass |  |  | Status: Incomplete
I/A/W MFR Procedure
Milliamp Process Clamp Meter
Backlight Test |  |  |  |  | N/A | Pass |  | Fail
Display Test |  |  |  |  | N/A | Pass |  | Fail
Keypad Test |  |  |  |  | N/A | Pass |  | Fail
Spotlight Test |  |  |  |  | N/A | Pass |  | Fail
DC Current Tests | [5 empty] | ± 0.2% of RDG + 5D
20.99 mA | 4 | mA |  |  | N/A | 3.94 | mA | 4.06 | mA | 0.06 |  |  | *Zero between each reading and
| -4 | mA |  |  | N/A | -4.06 | mA | -3.94 | mA | 0.06 |  |  | allow settling time before zeroing
| 12 | mA |  |  | N/A | 11.93 | mA | 12.07 | mA | 0.07
| -12 | mA |  |  | N/A | -12.07 | mA | -11.93 | mA | 0.07
| 20 | mA |  |  | N/A | 19.91 | mA | 20.09 | mA | 0.09
| -20 | mA |  |  | N/A | -20.09 | mA | -19.91 | mA | 0.09
± 1.0% of RDG + 5D
99.9 mA | 100 | mA |  |  | N/A | 98.5 | mA | 101.5 | mA | 1.5
| -100 | mA |  |  | N/A | -101.5 | mA | -98.5 | mA | 1.5
[10 empty] | Blank Cell Check | 12 | Fail

## Accredited
JM Test Systems, Inc. |  |  | Calibration Data Sheet |  |  | 0 | [6 empty] | 4 | Color coded formatting for t.u.r.
0 |  |  | 0 | [9 empty] | =FIXED(U15,2-(1+INT(LOG10(ABS(U15)))),TRUE)
0 | [12 empty] | This formula automatically rounds to two significant figures, regardless of how many decimal places
0 |  |  | 0 |  |  | 0
Function/Range |  |  | Out | Measured Value |  | Acceptance | Uncertainty | T.U.R.
Test Description |  | Nominal Value | ! | As-Found | As-Left | Limits | (K=2) | ( if < 4:1 )
TEST RESULT: |  |  |  | Pass | Pass | Status: Incomplete
I/A/W MFR Procedure
Milliamp Process Clamp Meter
Backlight Test |  |  |  |  | N/A | Pass/Fail
Display Test |  |  |  |  | N/A | Pass/Fail
Keypad Test |  |  |  |  | N/A | Pass/Fail
Spotlight Test |  |  |  |  | N/A | Pass/Fail |  |  |  |  | (3.3 to 33) mA | 0.018% + 0.25 µA
DC Current Tests | [5 empty] | ± 0.2% of RDG + 5D |  |  | no. of digits | tol± | Std. range | % rdg. | Rng. floor | limits± | k=1 | resol. | resol/sqrt(12) | r.s.s. | k=2 | rounded | t.u.r.
20.99 mA | 4 | mA |  |  | N/A | ±0.06 mA | 0.0059 mA |  | 2 | 0.06 | 33 mA | 0.018 | 0.00025 | 0.00097 | 0.000485 | 0.01 | 0.0028868 | 0.0029272 | 0.0058544 | 0.0059 | 10.1695
| -4 | mA |  |  | N/A | ±0.06 mA | 0.0059 mA |  | 2 | 0.06 | 33 mA | 0.018 | 0.00025 | 0.00097 | 0.000485 | 0.01 | 0.0028868 | 0.0029272 | 0.0058544 | 0.0059 | 10.1695
| 12 | mA |  |  | N/A | ±0.07 mA | 0.0063 mA |  | 2 | 0.07 | 33 mA | 0.018 | 0.00025 | 0.00241 | 0.001205 | 0.01 | 0.0028868 | 0.0031282 | 0.0062563 | 0.0063 | 11.1111
| -12 | mA |  |  | N/A | ±0.07 mA | 0.0063 mA |  | 2 | 0.07 | 33 mA | 0.018 | 0.00025 | 0.00241 | 0.001205 | 0.01 | 0.0028868 | 0.0031282 | 0.0062563 | 0.0063 | 11.1111
| 20 | mA |  |  | N/A | ±0.09 mA | 0.0069 mA |  | 2 | 0.09 | 33 mA | 0.018 | 0.00025 | 0.00385 | 0.001925 | 0.01 | 0.0028868 | 0.0034697 | 0.0069394 | 0.0069 | 13.0435
| -20 | mA |  |  | N/A | ±0.09 mA | 0.0069 mA |  | 2 | 0.09 | 33 mA | 0.018 | 0.00025 | 0.00385 | 0.001925 | 0.01 | 0.0028868 | 0.0034697 | 0.0069394 | 0.0069 | 13.0435
± 1.0% of RDG + 5D
99.9 mA | 100 | mA |  |  | N/A | ±1.5 mA | 0.061 mA |  | 1 | 1.5 | 330 mA | 0.018 | 0.0025 | 0.0205 | 0.01025 | 0.1 | 0.028868 | 0.030633 | 0.061266 | 0.061 | 24.5902
| -100 | mA |  |  | N/A | ±1.5 mA | 0.061 mA |  | 1 | 1.5 | 330 mA | 0.018 | 0.0025 | 0.0205 | 0.01025 | 0.1 | 0.028868 | 0.030633 | 0.061266 | 0.061 | 24.5902
End of Report | [10 empty] | (33 to 330) mA | 0.018% + 2.5 µA